## Slide 1
Optical Circulator
Analog-to-Digital Convertor
Photodetector
Demux
Mode-Locked
Fiber Laser
Dispersion Compensating
Fiber
EDFA
WDM Filter
1. Quantitative Phase Imaging
Dispersion Compensating
Fiber
Raman
Pump Lasers
Raman
Pump Lasers
WDM
WDM
2. Amplified Time-Stretch System
b. Phase Extraction
a. Pulse Synchronization
3. Big Data Analytics
Pixels
Frames
Wavefront
Nonlinear
Fiber
Fiber Coupler
Dispersive Gratings
Wave Plates
Mirror
Mirror
Beam Splitter
Target
Objective Lens
Spherical Lens
d. Machine Learning
λ1t1
λ1t2
λ2t1
λ3t1
λ1t3
…
…
c. Feature Extraction
FOV
Light
Flow
Notes: Fig 1. Time stretch quantitative phase imaging (TS-QPI) system simultaneously acquires brightfield images and quantitative phase images. Pulses from a mode-locked fiber laser (center wavelength = 1565 nm, repetition rate = 36.128 MHz, pulse width ~ 100 fs) are spectrally broadened with a highly nonlinear fiber to approximately 100 nm bandwidth. A dispersive fiber with an overall dispersion of 60 ps/nm is used to temporally broaden pulses to 1.2 ns, so an erbium doped fiber amplifier (EDFA) can amplify them without any distortion. Amplified pulses enter a coarse wavelength division multiplexing (WDM) filter, and the output of 1591 nm channel is used to shape laser pulses with a considerably flat spectrum over 1581 nm to 1601 nm bandwidth. These pulses pass through an optical circulator and are coupled to free-space with a fiber collimator.
Box 1 shows a train of broadband optical pulses spatially dispersed into a train of rainbows illuminating the target as line scans. The spatial features of the target are hence encoded into spectrum of broadband optical pulses each representing one 1-D frame. The ultra-short optical pulse illumination freezes the motion of cells during high speed flow to achieve blur-free imaging with a throughput of at least 100,000 cells/s (limited by the flow channel). The phase shift at each location within the field of view is embedded into the interference patterns using by a Michelson interferometer.
Box 2: The interferogram pulses were then stretched in time so that spatial information could be mapped into time through time-stretch dispersive Fourier transform (TS-DFT) and then captured by single pixel photodetector. The loss of resolution at high shutter speed is compensated by stimulated Raman amplification during time stretch.
Box 3: (1) Pulse synchronization: The envelope and phase of the time-domain signal was firstly mapped into series of spatial information , forming a scanning line brightfield image and phase contrast image, illuminated by the optical pulse at time . It was because within each optical pulse, the spatial information is one-to-one mapped into spectral domain and spectral domain is stretch in time , where is the relative group delay time of each frequency components within each pulse with regards to the central wavelength. These line-scan images based on were then cascaded into a two dimensional image corresponding to , where the second dimension is the spatial mapping of time elapse. Intensity measurement is calibrated based on the pulses during when cells are absent. (2) The biomass density of a cell leads to a spatially varying optical phase. When a rainbow linescan passes through the cells, the changes in refractive index at difference locations will cause phase walk-off at each interrogation wavelength. Hilbert Transform and unwrapping algorithm are used to generate the spatial information of phase shift. (3) Spatially dispersed rainbow creates multiple interrogation points for cell imaging. Decoding the phase shifts in each pulse at each wavelength and remapping into space reveals information about the distribution about protein concentration within cells. The transparency coefficient of each cell, corresponding to the fluctuation of intensity induced by the cell, is obtained from the amplitude of the slowly varying envelope on interferogram. Thus we achieve both image reconstruction of both brightfield and phase contrast simultaneously.
Fig 1. Time stretch quantitative phase imaging (TS-QPI) system simultaneously acquires brightfield images and quantitative phase images. Pulses from a mode-locked fiber laser (center wavelength = 1565 nm, repetition rate = 36.128 MHz, pulse width ~ 100 fs) are spectrally broadened with a highly nonlinear fiber to approximately 100 nm bandwidth. A dispersive fiber with an overall dispersion of 60 ps/nm is used to temporally broaden pulses to 1.2 ns, so an erbium doped fiber amplifier (EDFA) can amplify them without any distortion. Amplified pulses enter a coarse wavelength division multiplexing (WDM) filter, and the output of 1591 nm channel is used to shape laser pulses with a considerably flat spectrum over 1581 nm to 1601 nm bandwidth. These pulses pass through an optical circulator and are coupled to free-space with a fiber collimator.
Box 1 shows a train of broadband optical pulses spatially dispersed into a train of rainbows illuminating the target as line scans. The spatial features of the target are hence encoded into spectrum of broadband optical pulses each representing one 1-D frame. The ultra-short optical pulse illumination freezes the motion of cells during high speed flow to achieve blur-free imaging with a throughput of at least 100,000 cells/s (limited by the flow channel). The phase shift at each location within the field of view is embedded into the interference patterns using by a Michelson interferometer.
Box 2: The interferogram pulses were then stretched in time so that spatial information could be mapped into time through time-stretch dispersive Fourier transform (TS-DFT) and then captured by single pixel photodetector. The loss of resolution at high shutter speed is compensated by stimulated Raman amplification during time stretch.
Box 3: (1) Pulse synchronization: The envelope and phase of the time-domain signal 𝐼_𝑃𝐷 (𝑡_𝑖,𝑡_𝑑 ) was firstly mapped into series of spatial information 𝐼_𝑃𝐷 (𝑡_𝑖,𝑥_𝑗 ), forming a scanning line brightfield image and phase contrast image, illuminated by the optical pulse at time 𝑡_𝑖. It was because within each optical pulse, the spatial information is one-to-one mapped into spectral domain 𝑥_𝑗→𝜆_𝑗 and spectral domain is stretch in time 𝜆_𝑗→𝑡_𝑑^((𝑖)), where 𝑡_𝑑^((𝑖)) is the relative group delay time of each frequency components within each pulse with regards to the central wavelength. These line-scan images based on 𝐼_𝑃𝐷 (𝑡_1,𝑥), 𝐼_𝑃𝐷 (𝑡_2,𝑥), 𝐼_𝑃𝐷 (𝑡_3,𝑥)… were then cascaded into a two dimensional image corresponding to 𝐼_𝑃𝐷 (𝑦,𝑥), where the second dimension 𝑦 is the spatial mapping of time elapse. Intensity measurement is calibrated based on the pulses during when cells are absent. (2) The biomass density of a cell leads to a spatially varying optical phase. When a rainbow linescan passes through the cells, the changes in refractive index at difference locations will cause phase walk-off at each interrogation wavelength. Hilbert Transform and unwrapping algorithm are used to generate the spatial information of phase shift. (3) Spatially dispersed rainbow creates multiple interrogation points for cell imaging. Decoding the phase shifts in each pulse at each wavelength and remapping into space reveals information about the distribution about protein concentration within cells. The transparency coefficient of each cell, corresponding to the fluctuation of intensity induced by the cell, is obtained from the amplitude of the slowly varying envelope on interferogram. Thus we achieve both image reconstruction of both brightfield and phase contrast simultaneously.

## Slide 2
White Blood Cell Line (OT-II)
Optical Phase Image
Optical Loss Image
Colon Cancer Cell Line (SW-480)
Optical Phase Image
Optical Loss Image
Cell Flow
Cell Flow
100 µm
Cell Flow
Cell Flow
100 µm
100 µm
100 µm
Notes: Fig 2. phase image (upper) and optical loss image (lower) of SW480 (green) and OT-II cells (blue). The transparency of cells are affected by the attenuation of light going through the cells, as well as the scattering from the surface of the cells and from the internal cell organelles. The intensity is extracted from the lower frequency envelope variation of the waveform and relates to the transparency of individual cells in a flow stream. The phase information is extracted from the analytic forms of each pulse using Hilbert Transformation, followed by a phase unwrapping algorithm.
For image reconstruction, the time-domain signal was firstly mapped into series of spatial information , forming a scanning line image illuminated by the optical pulse at time . It was because within each optical pulse, the spatial information is one-to-one mapped into spectral domain and spectral domain is stretch in time , where is the relative group delay time of each frequency components within each pulse with regards to the central wavelength. These line-scan images were then cascaded into a two dimensional image , where the second dimension is the spatial mapping of time elapse.
Quantitative phase image (upper) and transparency image (lower) of SW480 (green) and OT-II cells (blue). The transparency of cells are affected by the attenuation of light going through the cells, as well as the scattering from the surface of the cells and from the internal cell organelles. The intensity is extracted from the lower frequency envelope variation of the waveform and relates to the transparency of individual cells in a flow stream. The phase information is extracted from the analytic forms of each pulse using Hilbert Transformation, followed by a phase unwrapping algorithm.
For image reconstruction, the time-domain signal 𝐼_𝑃𝐷 (𝑡_𝑖,𝑡_𝑑 ) was firstly mapped into series of spatial information 𝐼_𝑃𝐷 (𝑡_𝑖,𝑥_𝑗 ), forming a scanning line image illuminated by the optical pulse at time 𝑡_𝑖. It was because within each optical pulse, the spatial information is one-to-one mapped into spectral domain 𝑥_𝑗→𝜆_𝑗 and spectral domain is stretch in time 𝜆_𝑗→𝑡_𝑑^((𝑖)), where 𝑡_𝑑^((𝑖)) is the relative group delay time of each frequency components within each pulse with regards to the central wavelength. These line-scan images 𝐼_𝑃𝐷 (𝑡_1,𝑥), 𝐼_𝑃𝐷 (𝑡_2,𝑥), 𝐼_𝑃𝐷 (𝑡_3,𝑥)… were then cascaded into a two dimensional image 𝐼_𝑃𝐷 (𝑦,𝑥), where the second dimension 𝑦 is the spatial mapping of time elapse.

## Slide 3
Notes: Fig. 3 Three-dimensional scatter plot based on size, protein concentration, and transparency of the cells measured by Time Stretch Quantitative Phase Imaging. The green and blue dots are two-dimensional projections on the plane containing every combination of two parameters. The protein concentration corresponds to the local phase shift within the cell, which was converted into a local optical path difference mapping. The transparency is a parameter describing the intensity fluctuation caused by cell scattering and absorption.

## Slide 4
1
2
3
Circularity
Orientation
Diameter-RB
Diameter-FL
Tight Area
Loose Area
Perimeter
Major Axis
Median Radius
OPD-1
OPD-2
Refractive Index
Absorption-1
Absorption-2
Scattering-1
Scattering-2
Circularity
Orientation
Diameter-RB
Diameter-FL
Tight Area
Loose Area
Perimeter
Major Axis
Median Radius
OPD-1
OPD-2
Refractive Index
Absorption-1
Absorption-2
Scattering-1
Scattering-2
1
-1
0
Notes: Fig 4. Pairwise correlation plot visualized as heat map showing all major 16 features extracted from the images. Diagonal elements of the matrix are correlation of each parameter with itself, i.e. the autocorrelation. The set of parameter with the red color are highly correlated. The subset in the lower left show high correlation because they all measure morphological features. Also the subset in the upper right are correlated as they all measure attenuation and light scattering. This suggest that the dataset can be adequately represented by a smaller set that excludes highly correlated parameters.

## Slide 5
Notes: Fig 5. Ranking of feature performance based on the accuracy in 1D classification. The value is the area under ROC curve of each individual parameter.
Blue color shows morphological parameters includes diameter along the interrogation rainbow, diameter along the flow direction, tight cell area, loose cell area, perimeter, circularity, major axis length, orientation, and median radius. Orange bars show phase shift information include optical path length difference, refractive index difference and refractive index difference after noise reducing. Green bars show light scattering information.
Red box shows the best performed parameters in these three categories respectively.
Morphology contains most information. Others also have information.

## Slide 6
1-Specificity
Sensitivity
ROC
Feature Space
Decision
Making
Logistic Regression Chain
Image Fusion
d
d
Notes: Fig 6. Neural network architecture and training via Area Under the Curve (AUC) of the Receiver Operating Characteristics (ROC).
Multivariate features of each measurement (cell) are feed into the neural network and the output shows classification labels.
We train the weights of the last layer perceptron by optimizing the area under ROC curve. Each ROC curve corresponds to a set of weighting matrix to one of the output unit generated by scanning the bias weight. AUC is proven to be statistically consistent and more discriminating than accuracy.

## Slide 7
Protein Concentration
Attenuation
(d)
Size
Attenuation
(c)
Size
Protein Concentration
(b)
(a)
Notes: Fig. 8 Receiver Operating Characteristics (ROC) curves for binary classification of white blood cell (OTII) and cancer cells (SW480). The grey diagonal line shows results of random guess. Blue curves show the classifier performance using all 16 biophysical features extracted from the time-stretch QPI images.
We show ROC for classification based on single parameters. Red, green and orange curves show the classifier decision made using only the three major physical features: diameter, optical path length, and attenuation, respectively. Multivariate analysis based on time-stretch QPI shows significant improvement in classification. For each group we show the ROC for 10 fold cross validation data set. Additionally purple curves shows the classifier optimized with the first principal component in PCA feature space after data cleaning. Each kind of classifiers is shown with ten find curves demonstrating the results with ten-fold cross validation. Here we take 90% data for training and the other 10% for test. In the ten-fold cross validation, the total data set is split into 10 equal subsets and each time the test is performed for one part while the other nine parts are used for training. The fact that the classifiers remain almost unchanged during the ten runs of cross validation shows consistency and robustness of the data set.
Accuracy:
All 16 features averaging 10 folds after global cleaning: 0.9551
Size measurement averaging 10 folds after global cleaning: 0.8423
Size-RI measurement averaging 10 folds after global cleaning: 0.8420
Size measurement averaging 10 folds after local cleaning: 0.7806
Size-RI measurement averaging 10 folds after local cleaning: 0.8905
(1-0.7806)/(1-0.9551) = 4.8864
(0.9551-0.7806)/0.7806 = 22.4%

## Slide 8
Principal Component 1
Principal Component 2
Principal Component 3
Principal Component 4
Principal Component 5
Principal Component 6
Principal Component 7
Principal Component 8
Principal Component 9
Principal Component 10
Principal Component 11
Principal Component 12
Principal Component 13
Principal Component 14
Principal Component 15
Principal Component 16
Notes: Fig 7. (a-c) Decision boundary (red lines) in simultaneous two-parameter scatter plots using ten-fold cross validation. It shows the consistency and robustness of supervised classification. (a) diameter and protein concentration (b) diameter and scattering measurement (c) protein concentration and scattering.
(d) Performance ranking of each individual principal components in PCA space. The values of each bar show the percentage of the total variance explained by each principal components, accounting for the variability expressed in the data. Principal components with larger variability don’t necessarily give high accuracy in classification.
.

## Slide 9
(a)
(b)
Notes: Fig. 9 (a) Study of algae cells (Chlamydomonas reinhardtii) time-stretch QPI. Three-dimensional scatter plot based on size, protein concentration, and transparency of the cells measured by time-stretch QPI, with 2D projections for every combination of two parameters. Insert: Conventional flow cytometry using forward scattering and side scattering is not enough to distinguish the difference between high-lipid content and low-lipid content algae cells. Time-stretch QPI is much more effective in separating the two algae populations. (b) Receiver Operating Characteristics (ROC) curves for binary classification of normal and lipid-rich algae species. Blue curve shows the classifier performance using all 16 physical features extracted from the time-stretch QPI images. Red, green and orange curves show the classifier decision made using only the three major physical features, rainbow diameter, optical path length difference, and attenuation, respectively. Additionally we show ROC curves in purple for classification based on the first principal components in PCA space. The diagonal line shows results of random guess.

## Slide 10
sample
outlet
sheath
(a)
(b)
Notes: Appendex: microfluidic channel